W300PG&W300G-Bluetooth Broadcast Protocol

Data structure:

| Serial number | Length (bytes) | Meaning | Description |
| --- | --- | --- | --- |
| 0-1 | 2 bytes | Manufacturer ID | 0x5E5B |
| 2 | 1 byte | Power level + number of beacons | bit[0:3]: Number of positioning beacons scanned, bit[4:7]: 0 - 4, a total of 5 power levels |
| 3 | 1 byte | Status identifier | Bit0: Wearing status (1 indicates wearing, 0 indicates not wearing), Bit1: Network registration status (1 indicates network connection, 0 indicates network disconnection), Bit2: SOS status (1 indicates SOS mode, 0 indicates normal mode), Bit3: Sedentary status (1 indicates alarm mode, 0 indicates normal mode), Bit4: Fall status (1 indicates alarm mode, 0 indicates normal mode), Bit5~Bit7: Reserved and currently meaningless |
| 4 | 1 byte | Heart rate | Heart rate data |
| 5 | 1 byte | Systolic blood pressure | Blood pressure - systolic blood pressure data |
| 6 | 1 byte | Diastolic blood pressure | Blood Pressure - Diastolic Blood Pressure Data |
| 7 | 1 byte | Blood oxygen | Blood oxygen data |
| 8 | 2 bytes | Body temperature | Body temperature 10, big-endian representation, for example, if the body temperature is 36.8 degrees, then 36.810= 368, hexadecimal 0x0170 (in the data message) |
| 10-12 | 3 bytes | Step count | Big-endian representation, such as step count 3211, then expressed as 0x00 0x0C 0x8B |
| 13 | 1 byte | Bluetooth broadcast interval | Unit: seconds |
| 14 | 1 byte | Positioning interval | Bit7: 1 - Unit is minutes; 0 - Unit is seconds, Bit6~Bit0: Interval time |
| 15-26 | 12 bytes | Bluetooth Beacon 1 | 15-16: The minor of the scanned beacon 1, the default is 0 if not scanned, 17: The rssi of the scanned beacon 1, the signal values are all negative, take its positive part, if -50, then take 50 |
|  |  | Bluetooth Beacon 2 | 18-19: The minor of the scanned beacon 2, the default is 0 if not scanned, 20: The rssi of the scanned beacon 2, the signal values are all negative, take its positive part, if -50, then take 50 |
|  |  | Bluetooth Beacon 3 | 21-22: The minor of the scanned beacon 3, the default is 0 if not scanned, 23: The rssi of the scanned beacon 3, the signal values are all negative, take its positive part, if -50, then take 50” |
|  |  | Bluetooth Beacon 4 | 24-25: Minor of scanned beacon 4, default is 0 if not scanned, 26: RSSI of scanned beacon 4, signal values are negative, take its positive part, if -50, then take 50 |
| Total | 27 bytes |  |  |

Example:

Generally, for Bluetooth gateways delivered for docking, depending on different firmware versions, the location of this message is in adv_raw or srp_raw19FF5E5B310363744E6301700000270101735753

19FF: Start, followed by manufacturer data

5E5B: Manufacturer ID--5E5B

31: Battery level and number of beacons 16Hexadecimal to binary-->00110001 bit0-3：0001-->Convert to10decimal1--1beacon bit4-7:0011-->Convert to10decimal3--Battery level3

03: Status 16Hexadecimal to binary 00000011 Bit0:1--Wearing status, Bit1:1--Networked status has been networked, Bit2:0--sos status-SOS not triggered normal mode,

Bit3:0--Sedentary status-Sedentary normal mode not triggered, Bit4:0--Fall status-Fall normal mode not triggered

63: Heart rate 63-->16Hexadecimal to decimal-->99---Heart rate is99

74: Blood pressure-systolic blood pressure 74-->16Hexadecimal to decimal-->116---Blood pressure-systolic blood pressure is116

4E: Blood pressure-diastolic blood pressure 4E-->16Hexadecimal to decimal-->78---Blood pressure-diastolic blood pressure is78

63: Blood oxygen 63-->16Hexadecimal to decimal-->99---Blood oxygen is99

0170: Body temperature 0170-->16Hexadecimal to decimal-->368--->368*0.1-->Body temperature is36.8

000027: Number of steps 000027-->16Hexadecimal to decimal-->39---Number of steps is39

01: Broadcast interval 01-->16Hexadecimal to decimal-->1  Broadcast interval is1seconds

01: Positioning interval 01-->16Hexadecimal to binary-->00000001  bit7:0, the unit is seconds, Bit6~Bit0:0000001-->Converted to decimal:1. The positioning interval is1seconds

7357:minor Converted to big endian:5773-->16Hexadecimal to decimal-->22387---minor is22387

53: Signal value 53-->Convert to10decimal-->83-->Add a minus sign-->-83--Signal value is-83